## Slide 1: Compact OPO service Training
1st & 2nd March 2017

## Slide 2: Content
Introduction to COPO
Pump sources for COPO
Unpacking the COPO
COPO connections, Head and Panel PC
Software Interface
  Customer Interface
  Service Interface
Pre-Installation
COPO Alignment overview

## Slide 3: Chameleon compact opo
Closed box / hands free PP-Fan-OPO
Reduced size due to folded cavity
Temperature stabilized / purged
2 operation modes: Ti:Sa pass through / 15% Ti:Sa + OPO
Computer controlled phase-matching / Signal-tuning / power optimization
OPTICAL PARAMETRIC OSCILLATOR

## Slide 4: OPO operation
COPO is designed specifically to use the Chameleon as a pump source
Extends the WL coverage of the Chameleon into the near IR
Achieved by frequency conversion of the original WL into two longer WLS
λpump < λsignal < λidler
e.g 800nm pump < 1100nm signal < 2933nm idler
Fan-poled crystal – moves in vertical direction, different poling periods are available, allow phase-matching for a given combination of pump and signal wavelengths

## Slide 5: the COPO FAMILY
Standard Compact OPO
Compact OPO-Vis
Compact OPO Vis Output Options Intro
Chameleon-OPO-Vis-Data-Sheet
Chameleon MPX
Chameleon-MPX-Data-Sheet
Max IR
Idler (option)
Chameleon-OPO-Data-Sheet

## Slide 6: Pump sources for the Compact OPO
Chameleon Ultra 80MHz
Tuning range – 680nm to 1080nm
COPO pump range – 740nm to 880nm
Chameleon Vision 80MHz
Tuning range – 680nm to 1080nm
COPO pump range – 760nm to 880nm
*Systems not suitable for pumping Compact OPO
All Chameleon 90MHz systems - due to incompatibility of rep rate;
Vision-S

## Slide 7: Unpacking the Compact opo
The Compact OPO will be shipped with three boxes
Compact OPO head, Panel PC & Accessories
MRU
Chiller
(*CoolFlow IGE)

## Slide 8: Accessories
Cables
Mains cables
USB cable
RS-232 null modem cable (for connection with external PC)
Tools
Chiller hoses
Laser safety tube
Foot clamps (x3)
Fixing screws
Alignment apertures (x4)
Infrared sensor card
Gloves (x6)
Hair net (x2)
Hex keys for mirror adjustment (2mm & 2.5mm)

## Slide 9: COPO connections

## Slide 10: Panel pc

## Slide 11: Connections to Panel pc

## Slide 12: Software interface - user mode
Old SW
Current SW

## Slide 13: Service access (old S/W)
Service Interface
Service manual section 7.0
To enter service mode you need to change the target line. At the end of the line add;
space -- space service ( -- service)

## Slide 14: Service access (current S/W)
To enter service mode select ‘config’ tab, then type ‘expert=service’.
To exit service mode type ‘expert=0’.

## Slide 15: Custom tab

## Slide 16: Mirror tab
Starting Position for Mirrors is always centered

## Slide 17: Tuning tab

## Slide 18: Config 1 tab

## Slide 19: Config 2 tab

## Slide 20: Config 3 tab

## Slide 21: Pre-Installation
Chameleon laser must be installed as per standard service procedures
In particular CW and QS pump powers must have been calibrated using a Rees or Wavescan spectrometer
The laser must be thermally stabilised
The pump laser and OPO should be configured in-line and no intermediate steering optics
Recommended distance between laser and OPO is 60-150mm
Service engineer should carry a USB mouse

## Slide 22: COPO cavity

## Slide 23: Cavity
3 Sensor Elements
Spectrometer
Fast PD
Slow PD
Active Tuning Elements

## Slide 24: Alignment *Note this is a Brief overview only
OPO Head Alignment Alignment mode
Spectrometer & Fast PD Alignment 800nm
Pump In Alignment Alignment mode
Cavity Alignment 1st Pass 800nm
Cavity Alignment 2nd Pass 800nm
Start OPO Adjust cavity length to start OPO generation
Optimisation of power and tuning

## Slide 25: Locate the settings file
Old SW – C:\Program Files\opo compact
New SW – C:\ProgramData\APE\compact opo

## Slide 26: OPO head alignment
Set the Chameleon to Alignment mode
Place one of the slit aperture in front of BS1 and the other one before S3
Align the OPO on the table so that the beam passes through both apertures
Clamp the OPO onto the table to avoid shifting the OPO during the alignment process

## Slide 27: Spectrometer & Fast pd Alignment
Block the pump-in beam path into the OPO cavity
Tune the Chameleon to 800nm
Maximise the spectrometer peak at 800nm using BS2 and SM1

## Slide 28: Pump In Alignment
Set the Chameleon to Alignment mode
Place one of the hole apertures in front on PM and the other one before CM2
Use BS1 and PM to align the beam onto both apertures

## Slide 29: Cavity Alignment 1st Pass
Tune the Chameleon to 800nm
Place one of the hole apertures in the centre of the cavity (between M2 and M3)
Using CM2 only, centre the bright spot on the alignment hole

## Slide 30: Cavity Alignment 1st Pass con’t
Use the OC to adjust the back reflection on the other side of the hole aperture

## Slide 31: Cavity Alignment 2nd Pass
Place the hole aperture in front of M4
Centre the spot on the front side of the aperture using CM1
Adjust the back side reflection using M4

## Slide 32: Cavity alignment 2nd pass con’t
Place the hole apertures back into the centre of the cavity (between M2 and M3)
Re-adjust the back reflection using M4

## Slide 33: Start OPO
Once alignment has been optimised, remove alignment aperture
Select a wavelength e.g. 1100nm
Select the tuning tab and press ‘SET’ on the crystal position
Move the output coupler (cavity length) until a bright flash is seen
Fine tune the OPO generation using the cavity length adjustments

## Slide 34: Optimisation of power and tuning
Adjust BS3 and SM2 to maximise the signal peak on the spectrometer
Place an external power meter at the OPO signal output
Adjust M4 and OC iteratively to optimise the signal power
Perform data collection
PM

## Slide 35: Approved sample images