## Slide 1: Chameleon Service Training
Dieburg, 06.-10.06.2016
Display Menu

## Slide 2: Overview
Main Menu
Service Menu
Calibration

## Slide 3: Power Supply Control Software

## Slide 4
Software – MAIN MENU (1)
MAIN MENU
►System Diagnostic
System Shutdown
Alignment Mode
Initialise
Service
Displayed when in service mode
Main menu, displayed when menu/select
is pressed from the “Home” screen

## Slide 5
Software – MAIN MENU (2)
MAIN MENU
►System Diagnostic
System Shutdown
Alignment Mode
Initialise
Service
SYSTEM DIAGNOSTICS
►RS-232 Baud Rate
System Information
Fault Screen
RS-232 Setting Baud rate 19200, 8,N,1 – adjusted
by using the power adjust knob
System information: PZT positions, wavelength
Chameleon hours and software versions for both
the Chameleon and Verdi
Fault screen: displays current faults

## Slide 6
Software – MAIN MENU (3)
MAIN MENU
►System Diagnostic
System Shutdown
Alignment Mode
Initialise
Service
SYSTEM DIAGNOSTICS
►RS-232 Baud Rate
System Information
Fault Screen
RS-232 Setting Baud rate 19200, 8,N,1 – adjusted
by using the power adjust knob
System information: PZT positions, wavelength
Chameleon hours and software versions for both
the Chameleon and Verdi
Fault screen: displays current faults

## Slide 7
Software – MAIN MENU (4)
MAIN MENU
►System Diagnostic
System Shutdown
Alignment Mode
Initialise
Service
System Shut down
Press Select to continue
Press Exit to abort
Initiates system cool down prior to total
system shut down

## Slide 8
Software – MAIN MENU (5)
MAIN MENU
►System Diagnostic
System Shutdown
Alignment Mode
Initialise
Service
Alignment mode
EXIT when finished
Reduces laser output to a nominal low power in
order to facilitate alignment in applications where
the normal operating power is likely to cause damage

## Slide 9
Software – MAIN MENU (6)
MAIN MENU
►System Diagnostic
System Shutdown
Alignment Mode
Initialise
Service
Initialize
Press Select to continue
Press Exit to abort
Basically the same as “Recovery” with old style Chameleons …

## Slide 10
Software – SERVICE MENU (1)
MAIN MENU
►System Diagnostic
System Shutdown
Alignment Mode
Initialise
Service
SERVICE MENU
Calibration set up
Calibration
Home Stepper
Manual Control
Humidity
Settings
More straight forward, we will demonstrate
this in the lab

## Slide 11
Software – SERVICE MENU (2)
CALIBRATION SET-UP

## Slide 12
Software – SERVICE MENU (3)
CALIBRATION

## Slide 13
Software – SERVICE MENU (4)
HOME STEPPER

## Slide 14
Software – SERVICE MENU (5)
MANUAL CONTROL

## Slide 15
Software – SERVICE MENU (6)
HUMIDITY

## Slide 16
Software – SERVICE MENU (7)
SETTINGS