## Slide 1: Pointing
Beam Shift Measurement

## Slide 2: Customer’s Concern
Laser Beam drifts away from Target:

## Slide 3: Chameleon Specification Part 1
Chameleon V2 has a potential to “Point away”, if GDD is slightly misaligned.

## Slide 4: Pointing Measurement: Setup
Setup:
  Place wedged beam-splitter in front of Chameleon head
  You will have about 95% transmission and two reflecting beams.
  Choose the front reflection for the measurement, and block the other one
  Steer the beam with two HR mirrors i.e. in zig-zag across the available space on the table
  Place the Thorlabs detector as far away as possible.
  Glasgow factory recommends a distance of at least 3m.

## Slide 5: Pointing Measurement: Setup

## Slide 6: Pointing Measurement: Setup
Table Setup 1

## Slide 7: Pointing Measurement: Setup
Tool Box contain:
Thorlabs Detektor
(USB based)
3 Mirror Holder
Manual and Software CDs
2 HR Optic & 1 Wedged BS

## Slide 8: Actual Measurement Recommendation
Measurement:
  Put Beam into center of detector
  Start i.e. with 680, or 700nm.
  Tune 100nm in steps up to maximum
  Measure the shift of each step
  With a Vision, it would also make sense to check on different GDD values

## Slide 9: Pointing Measurement Result

## Slide 10: Pointing Measurement Result

## Slide 11: Pointing Measurement: Calculation of Pointing Angle
Results:
  Calculate the deviation angle (also called “Pointing”)
  tan(alpha) = opposite leg / adjacent leg
  Having very small angles, you can neglect calculating “tan”
  Pointing angle alpha [µrad] = Detector value [µm] / Length between Laser and Detector [m]
Compare with specification
  The interpretation of the released Pointing specification of Glasgow colleagues is like this:
  Pointing deviation from lowest to highest wavelength point.
  Spec is: <320µrad/400nm

## Slide 12: Pointing Measurement: Longterm Drift
Laser Beam drifts away from Target in time (hours/weeks):
Drift in Time is not specified
Installed Long-Term Setup @ Uni Greifswald
PowerMax Software on Internet accessable PC
Bomgar Remote Session for picking measurement results
All Eye Safe: Fully Tubed
Zeiss Service needs to be present during installation & de-installation

## Slide 13: Pointing Measurement: Longterm Drift
USB PowerMax Thermopile
Intensity Control

## Slide 14: Pointing Measurement: Signal to Noise
Thorlabs Detector Noise is can be 100µm average.

## Slide 15: Discussion: Published Specification, Part 2
Consideration of only two points:
l(min) against l(max) < 320µrad/400nm
Irritating Spec-Sheet description.
Pointing behavior in-between the two points is not considered.
1 Beamdiameter (1,2mm) shift in 3,8m distance is still in spec!
No Long-Term Drift specified
Our released specification does not meet requirements of Zeiss microscope application – Discussion ongoing.

## Slide 17: Long-Term Drift Measurement

## Slide 18: Old Pointing Data Chameleon XR

## Slide 19: Drift caused by external Optics: