## Slide 1: Chameleon Service Training
Dieburg, 06.-10.06.2016
System Installation

## Slide 2: Overview
Preparation
Equipment
Installation Procedure
Acceptance Test
Customer Training

## Slide 3: Preparation
Check, if
  Laser Crate, MRU, Chiller & CoolFlow Packages were delivered
    Check shock and tilt sensors
  Acclimatisation / bring in lab
  Lab ready?
    Optics table
    Electricity: number of wall plugs / multiple plug socket
    Laser safety installations
    Clean / dust-free floor

## Slide 4: Equipment
Personal Safety Glasses
Notebook, GUI installed and tested
REES!
Oscilloscope (recommendation: > 100MHz, Tektronix)
Powermeter(s) (LM-10 & LM-45 for Verdi calibration)
Fiber Scope
DVM
DC Current Clamp
IR Viewer / IR Card
Fast Photocell (Thorlabs DET10 A/M)

## Slide 5: Tools
Chameleon Pick-off Mirror, PN#1058622
Fiber Cleaning Supplies  Video
ESD Wrist Strap!
Cables
  USB Cable, Type A-A
  RS-232 Cable
  USB-RS232 Adapter
  BNC Cable
Standard Tools like allen wrenges, screw drivers, M6 screws,…

## Slide 6: Installation Procedure
Shipment Inspection
  Outside appearance, take pictures of damages, report to Coherent if needed
  Check if shipment is complete
Unpacking Head & PSU together
  Minimum of two people for lifting
  Apply back-gentle movements: 52kg
  Do not disconnect fibers from PSU!!
Connect all cables
  MRU cable: press hard until it clicks
  Chiller: Cooling flow direction important
Check lab voltage & compare with PSU fuses

## Slide 7: Installation Procedure
Check proper seating of „Fiber Cable Ferrule Nut“
PN# 0170-839-00
Ground yourself first.

## Slide 8: Installation Procedure
Install Measurement Equipment
Turn on Laser after Warm-up
Perform Wavelength Sweep
Record & Safe Datalog file
Record & Safe EEPROM data

## Slide 9: Installation Procedure
EEPROM Data / Calibration Table

## Slide 10: Installation Procedure
Tuning Calibration
During manufacture, Chameleon is calibrated at 41 equally spaced wavelength points across the tuning range.
System calibration should be valid for the lifetime of the laser.
Recalibration in the field should only be carried out following appropriate investigation of potential fault areas.
Re-calibration by trained service engineer only !

## Slide 11: Acceptance Test
Check proper pulse quality & power level across full tuning range
Record „power versus wavelength“ data in installation report
Compare with specification

## Slide 12: Customer Training & Closure
Inform customer about
  Laser beam properties & Laser safety
  Turn-On & Off Procedures
  required, regular maintenance
  If required, install GUI on customer PC
  Explain Remote Session
  On-site measurements and performance of this unit
Show customer proper filled installation report and let him sign
Send signed installation report to Coherent Tech Support Team

## Slide 13: Turn-On and Off Procedures

## Slide 14: Customer Maintenance

## Slide 15: Still some time and energy ???

## Slide 16: Inform customer about:
Service Strategy &
LMC Advantages

## Slide 17: Technical Support Hotline
All technical questions can be applied to our Inhouse Support Team
Phone +49 6071 968-225
E-Mail service.dieburg@coherent.com
Or to you: Leave your business card on the optics table ;-)

## Slide 18: In case of an issue: Our Service Strategy
Within 12 months after installation
  + 1 month integration time
  Warranty
Beyond 12 months:
  Customer buys warranty extention, or
  Paid Service: Service visit & parts will be charged on event basis
    NO ARU program!
    Depot Repair costs are between 6,5k€ (minor) – 32,5k€ (Verdi repl.), or
    Even higher, if several issues have been found
    Full Depot Repair Time approx. 6-10 weeks

## Slide 19: Service Strategy
Warranty Extention
  Name: Productivity Plus (P-Plus)
  Free of charge Remote Session per Bomgar / TeamViewer
  Parts & Field Visit Costs inclusive
  Preferred Reaction Time
  Minimum Downtime
  ARU Program
  Important Requirement:
    Laser has to be in specification before it can join P-Plus.

## Slide 20: Service Strategy: P-Plus – Flyer 1/2

## Slide 21: Service Strategy: P-Plus – Flyer 2/2

## Slide 22: Service Strategy / P-Plus Matrix
 | Chameleon |  | 
System Lifetime Period | Warranty | Out of Warranty | Service Contract
 | 12 (24) months | > 12 (24) months | 
Plan | Option |  | 
Depot Repair Repaired system returned typically 4-6 weeks after receipt at Coherent | Not Applicable | Standard (Paid) | PPlus Bronze
Return Replacement Replacement system ships within 15 business days | Not Applicable | Not Applicable | PPlus Silver
Advance Replacement Ship from GLS within 2 working days (for 90% of all events) after customer informs of failed system | Standard (Free of Charge) | Not Applicable | PPlus Gold

## Slide 23: Inform customer about LMC: Laser Measurement & Control short: Powermeters
LMC Price Advantages
Leave LMC Flyer &
Send LMC catalogue
  Or give Link to Coherent website:
  https://www.coherent.com/products/?1940/Laser-Measurement-and-Control-Literature

## Slide 24: Now, let’s do the first Quiz…

## Slide 25: Some Info's about the daily Quiz
Altogether four quizzes of the course
To pass a quiz, you need > 60%
To get an A+, you need 100%
With bonus questions, you can achieve >100%
These questions are sophisticated, a little mean, but always related to Chameleon ;-)
Free internet allowed, but please try to find the answers yourself.