QSMART – INTEGRAGRATION OF AN EMPTY PUMP CHAMBER

PROCEDURE

Toolkit:

Metric Allen key

Special tool RM100004

Target (at 43mm)

System: QSMART

| SM010341 SM010342 SM010343 SM010344; Q-SMART850 PUMP CH W OPT S/E Q-SMART850 PUMP CH W/O OPT Q-SMART850 ROD ASSY Q-SMART850 ROD ASSY STD EXCH. |  |
| --- | --- |

Purpose: This document details how to change the content of the pump chamber

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |

# ESTABLISHMENT OF Q-SMART850 ROD ASSY

| 2.00 | 3/E10/0659 | O-Ring D15 C2 |
| --- | --- | --- |
| 2.00 | 3/PLA/1238 | Rod retainer |
| 8.00 | V/CHC/M0306A4 | Screw CHC M3X6 |
| 1.00 | RM100004 | Special tool rm100004 |

# DISASSEMBLING

- At this step, the pump chamber is outside the Laser head.

- (Refer to the procedure: “LH – SM010340 Pumping Chamber replacement”

- Remove the flashlamp cover by unscrew (in a cross) the 8 screws.

- Remove all the mechanical parts which maintain the rod.

- Remove the rod.

# REASSEMBLING AND ALIGNMENT

- Put the rod assy and centered it inside the pump chamber.

- Add the two O-rings (3/E10/0659)

- Add the two rod retainer (3/PLA/1238) and the 8 screws CHC 3x8 without screwing completely the screws.

Check the good alignment of the He-Ne (height=43mm): the pin hole must be in place.

You have to check with a target near field (adjust M1) and far field (adjust M2)

M1

Pin hole

He-Ne

M2

Target

- Install the pump chamber on the special tool RM100004. Install this two parts on the He-Ne way.

RM100004

Pumping chamber on his alignment tool (view of the top)

RM100004

HeNe incident    ___________

HeNe reflected   _ _ _ _ _ _ _ _

- When the adjustment of the reflected beam is done, tight in cross the 4 screws of rod retainers (3/PLA/1299) on both side. Check, that the rod hasn’t move when you have finished to tight.

- Reassembly of the pump chamber

| 1.00 | 3/E10/0667 | O-Ring D153 C2 |
| --- | --- | --- |
| 2.00 | 3/PLA/1299 | SEALING DUST RING |
| 2.00 | 3/ATC/0010 | Cover label + |
| 2.00 | 3/ATC/0013 | Cover label - |
| 1.00 | 3/ATC/0121 | Cover label HT |
| 8.00 | V/CHC/M0410A4 | Screws CHC M4x10 |

- Place the flashlamp cover and screw in the indicated order the 8 screws (CHC 4x10)

- Install and screw the 2 sealing dust ring (3/PLA/1229), and add the 2 Orings.