QSMART – LED CCA REPLACEMENT

PROCEDURE

Toolkit:

Allen keys

System: QSMART

| SM010351; Q-SMART LED CCA. |  |  |  |  |
| --- | --- | --- | --- | --- |

Purpose: This document details how to change the LED Card

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |
| revB | November 28, 2014 | CCA LED adds harness and require procedure modification |

# DISASSEMBLING

- Turn the system off

- Disconnect cables and water hoses from the side of the laser head

- Remove the laser head cover.

- Remove the side plate (7 BHC 3x6)

- Remove the two UNC 4-40 long ¼ 3/E13/0601 screws on the front panel

Connector attachment face avant

- Unsolder yellow and brown wires of the LED card on external contacts of the switch.

Transistor

- Disconnect the LED cable (J3) from the LB card

- Unscrew the screws holding the LB card (5 screws CHC3X6 + 4 MU3 washers).

- Remove temperature sensor, refer to “LH - SM010350- TEMP SENSOR ASSY”.

- Extract 2 temperature sensor wires from white connector (see photo)

- Remove the manual shutter 3/PLA1294 by unscrewing the screw CHC 3x6.

- Remove the 4/300/1057 LED card by unscrewing the 2 screws CHC 3x5 +2 washers

# IMPLEMENTATION OF THE LED CARD

- Insert the LED Card 4/300/1057 (you may have to deburr the edges) and fasten it (2 CHC 3x5 screws +2 ZU3 washers).

- Make sure that the axle pin rotates correctly.

- Insert 2 temperature sensor wires into the white connector.

- Fasten the manual shutter 3/PLA1294 (1 CHC 3x6screw)

- Check the proper functioning of the shutter.

- Pass the LED card wires below the LB card 4/300/1036. (refer to the picture)

Transistor

- Fasten with 4 CHC 3x8 screws + 4 M3U washers without tightening

- Connect the LED cable on the LB card (J3).

- Install the LB card connector on the front and tighten each screw.

- Fasten the transistor to the card with a CHC 3x6 screw + ZU3 washer.

- Weld the three pins of the transistor.

- Insert back the temperature sensor on head casting

Wires

- Weld the yellow and brown wires of the LED card on external contacts of the switch.

- Screw the 2 4-40 UNC long 3/E13/0601 ¼ screws on the front panel

Connector attachment face avant

- Close the Laser head and put warranty stickers 
Refer to "LH-Warranty Labels"