QSMART – CPU BOARD REPLACEMENT

PROCEDURE

Toolkit:

Allen keys

System	: QSMART

| 4/300/1077; Q-SMART KEY CCA. |  |  |  |  |
| --- | --- | --- | --- | --- |

Purpose: This document details how to change the CPU board

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |

# DISASSEMBLY

- Unplug the system and disconnect all cables

- Remove the CG cap

3/E13/0600

- Remove the 4 screws 3/E13/0600 (Rear panel)

3/E13/0600

Enlever les 4 Vis 3/E13/0600 (face avant). Enlever les deux flasques 3/PNT/0511 et 3/PNT/0512

- Remove the 4 screws 3/E13/0600 (front). Remove the two side panel

Disconnect the cable kit 4/300/1070 from the CPU card. Remove Tyraps.

- Disconnect the key card from the CPU as pictured.

Remove all BNC connections in the front

Disconnect the BNC security.

Remove the CPU card by unscrewing the 4 screws CHC 3x8 + 4 washers LU3.

# Integration of the CPU Board

| 4.00 | V/CHC-M0308-A4 | CHC 3x8 SCREW |
| --- | --- | --- |
| 4.00 | R/LU--D03--A2 | LU3 WASHER |

Fasten the CPU card using 4 screws CHC 3x8 + 4 washers LU3.

Check the BNC connections, it should connect well without straining.

Connect the key card to the CPU as pictured.

Tyraps

Connect the cable kit 4/300/1070 to the CPU card and attach each cable to the frame with tyraps as shown on the picture.

Connecter le kit câble 4/300/1070 à la carte CPU et au groupe de froid  et attacher chaque câble  au bâti avec des tyraps comme sur photo ci-contre.

# COVERING

| 8.00 | 3/E13/0600 | BLACK BHC 4x12 SCREW |
| --- | --- | --- |

Implement both side panels and fasten with 8 BHC Black 4x12 (3/E13/0600) screws

3/E13/0600

3/E13/0600

CAUTION: Leave a space between two covers

Reinstall the CG cap