QSMART –HEAT EXCHANGER AND FAN REPLACEMENT

PROCEDURE

Toolkit:

Allen keys

System	: QSMART

| 3/E03/0022A | Q-SMART CG HEAT EXCHANGER |
| --- | --- |
| 4/VEN/0046 | Q-SMART CG FAN |

Purpose: This document details how to change fans and heat exchanger

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |

# DISASSEMBLY

- Open the ICE and the cooling group

See procedure "CG - SM010355-cooling unit"

# IMPLEMENTATION of the CG FAN

Keep all screws and washers:

| 12.00 | V/CHC-M0412A2 | CHC M4-12 STAINLESS STEEL A2 |
| --- | --- | --- |
| 12.00 | R/MU--D04--A2 | MU4 WASHER |

Wires output

Air flow direction

Tube mounting instructions :

Lower tube

Upper tube

Fan wires to be placed here. Cut wires 40 mm outside heat exchanger. Remove insulation over 8mm

40mm

- Put together all wires with same colour. Insert in one connecting port. One port for one colour (blues with pink)

# Heat Exchanger

| 3.00 | 3/E13/0600 | BHC M4 BLACK SCREW |
| --- | --- | --- |

# WIRING

- Connect the cables that have been disconnected

Refer to "CG-Wiring"

# TEST AND CLOSING

- Perform the test of CG and close.

- Close the ICE

Refer to «CG-Test and closing" and "CG - SM010355-cooling group"