QSMART – COOLING GROUP REPLACEMENT

PROCEDURE

Toolkit:

Allen keys

System	: QSMART

| SM010355; Q-SMART COOLING GROUP ASSEMBLY. |  |  |  |  |
| --- | --- | --- | --- | --- |

Purpose: This document details how to change the COOLING GROUP

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |

# DISASSEMBLY

## Unplug the system

- Remove the cap

## Unscrew the four screws, two at front and two on the rear side as indicated in the pictures

## Slide back the panel as indicated in the picture

- Unscrew the 3 screws black 4x12 BHC (3/E13/0600) to remove the CG

- Disconnect the cables from the CG

3/E13/0600

- Open CG by removing the locking plate (9 BHC M3x10)

Unscrew the cap on the tank

Open CG by removing the locking plate (9 BHC M3x10)

Note: these 3 black BHC screws fasten the heater

# INSTALLATION OF COOLING GROUP

| 3.00 | 3/E13/0600 | BLACK BHC 4x12 SCREW |
| --- | --- | --- |

Fasten the CG to the chassis with 3 Black BHC 4x12 screws (3/E13/0600).

3/E13/0600

# CLOSING

| 8.00 | 3/E13/0600 | BLACK BHC 4x12 SCREW |
| --- | --- | --- |

3/E13/0600

Implement both side panel and fasten with 8 BHC Black 4x12 (3/E13/0600) screws

3/E13/0600

CAUTION: Leave a space between two covers

Reinstall the CG cap