QSMART – PUMP CHAMBER REPLACEMENT

PROCEDURE

Toolkit:

Metric Allen Keys

Special tool for interlock cover head RM100016A

Special tool SM020006A for software update

Ethernet cable

Powermeter

Scope and photodiode

Burn paper

System: QSMART

| SM010340 SM010341; PUMP CHAMBER W OPT PUMP CHAMBER W OPT S/E. |  |  |  |  |
| --- | --- | --- | --- | --- |

Purpose: This procedure describes the integration of the complete pump chamber of a QSmart

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |

# DISASSEMBLING

- Turn the system off

- Disconnect cables and water hoses from the side of the laser head side

- Remove the laser head cover.

- Remove the side plate (7 BHC 3x6)

- Disconnect the flash lamp cables on the TRIGGER card

- Unscrew the 2 O-rings located on both sides of the pump chamber

- Remove the screw which maintain the cable coming from the TRIGGER card on the pump chamber

- Remove the pump chamber ( unscrew the 4 screws )

3/E10/0666 O-rings

3/E10/0666 O-rings

3/E10/0353

O-rings on the water main

3/E10/0353

O-rings on the water main

# Installation of the pump chamber

- Place the optical head on an optical table

- Add 3 mounting feet in order to fix the Laser head to an optical table

- Place two O-rings D20 C4 (3/E10/0666) at the extremity of the pump chamber and 2 others rings (3/E10/0353) on the water main.

Be careful, the highest screws are the shortest screws and the lowest screws are the longest.

- Adjust the O-rings of « sealing dust », put in touch the O-rings with the body

3/E10/0666

3/E10/0666

- Plug the lamp flash cable on the TRI board

- Screw the cable of the TRI board on the pump chamber (CHC 3x6 + ZU3 washer)

The pump chamber is now re-assembled

# ADJUSTMENT OF THE LASER (Spiricon or Burn Paper)

Refer to Laser Adjustment procedure.

- Close the Laser head and put warranty stickers 
Refer to "LH-Warranty Labels"