QSMART – PUMP CHAMBER REPLACEMENT

PROCEDURE

Toolkit:

Allen keys set - metric

Cover interlock holder RM100016A

Software loading harness SM020006A

Optical table

System: QSMART

| SM010340 SM010341; PUMP CHAMBER W OPT PUMP CHAMBER W OPT S/E. |  |  |  |  |
| --- | --- | --- | --- | --- |

Purpose: This procedure describes the integration of the complete pump chamber of a QSmart

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |

# DISASSEMBLING

- Turn the system off

- Disconnect cables and water hoses on the laser head side

- Remove the laser head cover.

- Remove the side plate (7 BHC 3x6)

- Remove the trigger cable (keep the CHC 3x6 screw + ZU3 washer).

- Disconnect flash lamp  wires from Trigger Board

- Unscrew the 2 O-rings located on both sides of the pump chamber

- Remove the screw which maintain the cable coming from the TRI board on the pump chamber.

- Remove the pump chamber by unscrew the 4 screws

# IMPLEMENTATION OF THE NEW PUMPING CHAMBER

- Fasten the laser head on the optical table (3 points)

- Put the new sealing rings (3/E/0666) on each side then slide the pumping chamber on water pipes. First screw the 2 upper CHC 4x25 screws, then the lower CHC 4x40 ones

- Adjust nuts to push sealing rings on the main structure.

- Connect flash lamp wires on the trigger board

- Fasten the trigger cable (CHC 3x6 screw + ZU3 washer).

- Connect the power supply cable to the laser head

- Fill the cooling group water tank, verify water level (see User Manual)

- Connect water hoses. Check for color codes.

- Defeat cover interlock with tool RM100016A

- Turn on the power supply cabinet.

- Check that the Shutter Led lights up at turn on then shuts off

- Adjust tank water level

# LASER ADJUSTMENT (Beam analyzer or burn paper)

Refer to Laser Adjustment procedure.

- Close the Laser head and put warranty stickers 
Refer to "LH-Warranty Labels"