QSMART – POCKELS CELL REPLACEMENT

PROCEDURE

Toolkit:

Screwdrivers

Allen keys

Cover interlock holder RM100016A

Energy meter

Photodiode

Analyzer / polarizer 1/OUT/0019

Upgrade soft Q-Smart tool SM020006A

System: QSMART

| SM010348A; Q-SMART POCKELS CELL ASSY. |  |  |  |  |
| --- | --- | --- | --- | --- |

Purpose: This document details how to change the pockels cell

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |

# 1 DISASSEMBLING

- Turn the system off

- Disconnect cables and water hoses from the side of the laser head

- Remove the laser head cover.

- Remove the side plate (7 BHC 3x6)

- Remove the ground wire (1), remove the power connector on the card Q Switch (2) and remove the connectors on the pockels (3)

Position1 et 2

Position1 et 2

Position3

Position3

- Remove the unit from the foundry (3 screws)

Remove the quarter wave plate mount (and cylindrical lens if present)

- Remove the Pockels cell (1 CHC 4x30)

# 2 REASSEMBLING AND ALIGNMENT

M1

M1

Pin hole

Pin hole

Source He-Ne

Source He-Ne

M2

M2

- Check the good alignment of the He-Ne (height=43mm): the pin hole must be in place.

- You have to check with a target near field (adjust M1) and far field (adjust M2)

- Fasten the Pockels cell with CHC screws 4x30.

- Put the Analyzer/Polarizer 1/OUT/0019 on each side of the Pockels like shown.

Mire

Polarizer r*

Rod

Pockels cell

Analyzer

10 cm

Mire

Polarizer r*

Rod

Pockels cell

Analyzer

10 cm

- You should see a Maltese cross when the analyzer and the polarizer are perpendicular.

- Adjust the two axes of the support so the red dot (reference beam) passes through the center of the cross.

- After the adjustment, tighten the screws to block the movement of the pockels.

- Replace pockels housing with 3 screws and connect the power wires on the cell.

| 1.00 | V/BHC-M0306A4 | BHC 3X6 SCREW |
| --- | --- | --- |
| 1.00 | R-DEC-D03-- | LOCKING 3 WASHER |

- Fasten the unit to the foundry (3 screws BHC4x10 + 4 ZU4 washers)

- Replace the ground wire (1), connect the power connector on the card Q Switch (2) and the connectors on the pockels (3)

Position1 et 2

Position1 et 2

Position3

Position3

# 3 LASER ADJUSTMENT

Refer to the Laser Adjustment procedure

- Close the Laser head and put warranty stickers 
Refer to "LH-Warranty Labels"