QSMART – TEMP SENSOR ASSY REPLACEMENT

PROCEDURE

Toolkit:

Allen keys

System: QSMART

| SM010350; Q-SMART TEMP SENSOR ASSY.. |  |
| --- | --- |

Purpose: This document details how to change and adjust the laser head temperature sensor

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |

# DISASSEMBLING OF THE SENSOR

- Turn the system off

- Disconnect cables and water hoses from the side of the laser head

- Remove the laser head cover.

- Remove the side plate (7 BHC 3x6)

- Unscrew (CHC 2.5x6) and remove the sensor support (3/MEC/4929)

- Unscrew the screws holding the LB card (5 screws CHC3X6).

Remove the 2 sensor wires

- Remove the sensor wires from the connector

- Cut the tape and remove the sensor (support + sensor + wires)

# IMPLEMENTATION OF SENSOR

| 1.00 | V/CHC-M2506A4 | CHC2.5X6 SCREW |
| --- | --- | --- |
| 5.00 | V/CHC-M0310A4 | CHC3X6 SCREW |
| 4.00 | R-MU--D03--A2 | MU3 WASHER |

- Insert the sensor wires in the connector like shown (2 last pins) and plug back all connectors on the card.

- Link between them the wires using adhesive tape.

- Pass the cables under the LB card 4/300/1036.

- Fasten with 5 screws CHC 3x6 + 4 washers

- Fasten the sensor SM010350 and its support with a CHC 2.5x6 screw.

- Reinstall the side plate and the laser head cover

- Close the Laser head and put warranty stickers 
Refer to "LH-Warranty Labels"