QSMART – LASER ADJUSTMENT

PROCEDURE

Toolkit:

Allen keys

Energy meter

Photodiode

Oscilloscope

System	: QSMART

Purpose: This document details how to change align the laser components

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |
| Rev. B | May 7, 2015 | Add free running checking |

# ADJUSTMENT OF THE LASER (Spiricon or Burn Paper)

- Connect the Laser head to the power supply

- Connect the coolant hoses (respect the color)

- Switch on the power supply

- Check that the Shutter LED on the head lights up then lights out.

- Adjust the level of water

- Set up the Special tool for interlock cover head RM100016A

- Disconnect the Qtouch and plug the RJ45 ethernet cable

- Launch the Qsmart-Service.exe application on your computer

Qsmart-Service application

- If there is any communication between the Qsmart-Service application on your computer and the power supply, adjust the parameter of IP address on your lab top. Enter 169.254.0.10

- Open the file : Modulo.ini and enter the IP 169.250.0.1

Now, the communication is establish

Let us return on the Laser head adjustment

### Remove the ¼ wave plate (which permits holdoff)

- Set and Check :

    - Flash voltage: 600V

    - Charger set point  160µs

    - Shutter head is open

- Start the flashlamp ONLY

- Align the rear mirror in order to have a round shape on the CCD Spiricon camera or on a burn paper.

- Optimize the output energy by turning the second ¼  wave plate located at the front of the Laser head

- Adjust the energy to 850mJ by increasing the Flashlamp Voltage «Charger set point»

- Remove the power meter and check the beam again

- Stop the flash lamp

- Re-install the hold off ¼ wave plate

- Start the flash lamp

- Plug a photodiode to a scope in order to look the flash lamp fluorescence

- Rotate the ¼wave plate to establish hold off

- If there is still free running, adjust the Pockels cell

- Go to : Laser Control tab

- Note the original value of « Charger Max », change the value to 680 V.

- Adjust the voltage incrementally up to 680 Volts.

Then, adjust the QSwitch incrementally up to 180 μs to make sure the cavity is closed.

- Find the maximum holdoff location and tighten the ¼ 	wave plate and Pockels cell

See next page for free running checking

When closing the cavity you are avoiding free running into the laser:

How to check if you have free running:

Step 1

Cut the beam with a beam pad to have the scatted the light and place a photodiode that see the beam way

FL ON

QS OFF

If you have a Gaussian curve = no free running, only the flashlamp discharge.

If you have a peak on the top of it, you have free running.

- Check again that there isn’t free running

The cavity is close….Stop the flashlamp

- Set :

    - QSwitch voltage : 160V

    - Charger set point : 600V

    - Flash pulse width 170µs

    - QS delay = 7µs

    - « Charger max » (enter the original value that you have note before)

### Optimize the alignment in Qswitch mode (rear mirror)

### Increase gradually the flashlamp voltage to obtain 850mJ

- Place the white teflon shield part.

Wait until this protection part is at the same temperature than the Laser head before screwing it

- Check the profile beam again

# SAVE THE PARAMETERS

- Save parameters      (Nota : Cmd = FPGM  two times)

- Shut down and start again the power supply in order to check the parameters are save in the LB board

# FINISHING WORKS

- Shut down the power supply

- Disconnect the i/o power cable and coolant hoses

- Set up the cover and screw it

- Re-connect the i/o power cable and coolant hoses

- Check the energy of the Laser

- Check the beam profile on a burn paper at 30cm, 1m and 2m