QSMART – KEY CARD REPLACEMENT

PROCEDURE

Toolkit:

Allen keys

System	: QSMART

| 4/300/1077; Q-SMART KEY CCA. |  |  |  |  |
| --- | --- | --- | --- | --- |

Purpose: This document details how to change the key card

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |

# DISASSEMBLING

- Unplug the system and disconnect all cables

- Remove the CG cap

3/E13/0600

- Remove the 4 screws 3/E13/0600 (Rear panel)

3/E13/0600

Enlever les 4 Vis 3/E13/0600 (face avant). Enlever les deux flasques 3/PNT/0511 et 3/PNT/0512

- Remove the 4 screws 3/E13/0600 (front). Remove the two side panels.

3/E13/0600

- Remove the CG (3 Black BHC 4x12 (3/E13/0600)).

- Disconnect the key card from the CPU as pictured.

- Remove the front panel ICE.

BHC4x8

BHC 4x6+LLU4

BHC 4x8

BHC 4x8+LLU3

- Remove the key card 4/300/1077 (3 screws CHC 3x6 + 3 washers CS3).

CHC 3x6 + CS3

# KEY CARD INTEGRATION

| 3.00 | V/CHC-M0306A4 | CHC 3x6 SCREW |
| --- | --- | --- |
| 3.00 | R/CS--D03--ZB | CS3 WASHER |

CHC 3x6 + CS3

- Fasten the key card 4/300/1077 (3 CHC 3x6 +3 CS3).

Be careful to put the key in front of the 0 when tightening.

- Attach the ICE front panel.

| 3.00 | V/BHC-M0406-A2 | BHC 4X6 SCREW |
| --- | --- | --- |
| 4.00 | V/BHC-M0408-A2 | BHC 4X8 SCREW |
| 3.00 | R/LLU-D04--A2 | LLU4 WASHER |

BHC4x8

BHC 4x6+LLU4

BHC 4x8

BHC 4x8+LLU3

Connect the key card to the CPU as pictured.

# COVERING

| 8.00 | 3/E13/0600 | BLACK BHC 4x12 SCREW |
| --- | --- | --- |

Implement both side panels and fasten with 8 BHC Black 4x12 (3/E13/0600) screws

3/E13/0600

3/E13/0600

CAUTION: Leave a space between two covers

Put back in place the CG cap