QSMART – POWER SUPPLY REPLACEMENT

PROCEDURE

Toolkit:

Allen keys

System	: QSMART

| SM010358; Q-SMART POWER SUPPLY. |  |  |  |  |
| --- | --- | --- | --- | --- |

Purpose: This document details how to change the power supply

| Revision | date | modification |
| --- | --- | --- |

# DISASSEMBLY

- Unplug the system and disconnect all cables

- Remove the CG cap

3/E13/0600

- Remove the 4 screws 3/E13/0600 (Rear panel)

3/E13/0600

- Remove the 4 screws 3/E13/0600 (front). Remove the two side panels

- Remove the 3 screws BHC 4x12 (3/E13/0600) from the PB

3/E13/0600

- Disconnect the cable kit 4/300/1070 from the PB card.

Remove the two connectors.

Remove the power supply

# INTEGRATION OF THE POWER SUPPLY

| 3.00 | 3/E13/0600 | BLACK BHC 4x12 SCREW |
| --- | --- | --- |
| 2.00 | 4/CBL/0005 | SMALL TYRAP |

- Set up the two connectors as pictured

Put back in place the PB card into the ICE

Connect the cable kit 4/300/1070 to the PB card like shown.

Attach the cables with tyraps

Fasten the power supply with 3 BHC 4x12 (3/E13/0600)

3/E13/0600

# COVERING

| 8.00 | 3/E13/0600 | VIS BHC 4x12 NOIRE |
| --- | --- | --- |

Implement both side panel and fasten with 8 BHC Black 4x12 (3/E13/0600) screws

3/E13/0600

3/E13/0600

CAUTION: Leave a space between two covers

Put back in place the CG cap