QSMART – TRIGGER CARD REPLACEMENT

PROCEDURE

Toolkit:

Allen Keys

System: QSMART

| SM010352; Q-SMART TRIGGER CCA ASSY. |  |  |  |  |
| --- | --- | --- | --- | --- |

Purpose: This document details how to change the TRIGGER Card

| Revision | date | modification |
| --- | --- | --- |
| Initial issue | July 4, 2014 |  |

# DISASSEMBLING

- Turn the system off

- Disconnect cables and water hoses from the side of the laser head

- Remove the laser head cover.

- Unscrew the CHC 3x6 screw + ZU3 washer holding the trigger card wire

- Disconnect the flash cables from the Trigger card.

- Unscrew the 2 4x10 HCC screws holding the unit

- Remove the TRIGGER card 4/300/1056 from its 3/PLA/1267 support (3 CHC M3x8 screws + 3 ZU3 washers)

# IMPLEMENTATION OF THE TRIGGER CARD

| 1.00 | 4/300/1056 | TRIGGER CARD MOD EQUIPED |
| --- | --- | --- |
| 3.00 | V/CHC-M0308A4 | CHC3X8 SCREW |
| 2.00 | V/CHC-M0410A4 | CHC4X10 SCREW |
| 3.00 | R-ZU--D03--A2 | ZU3 WASHER |

- Fasten the TRIGGER card 4/300/1056 to its 3/PLA/1267 support (3 CHC M3x8 screws + 3 ZU3 washers)

- Fasten the unit to the foundry (2 4x10 HCC screws)

Connect the wires as shown

Connect the wires as shown

- Connect the flash wires like shown.

- Connect the flash cables to the TRIGGER card

- Fasten the TRGGER card wire like shown (CHC 3x6 screw + ZU3 washer)

- Close the Laser head and put warranty stickers 
Refer to "LH-Warranty Labels"