## Slide 1: Chameleon Service Training
Dieburg, 06.-10.06.2016
Precompensation
Vision Family

## Slide 2: Overview
Theory
Carpe Measurement
Setting the GDD Curve
Chameleon Vision – Precomp Module

## Slide 3: Overview
The Vision-Family features an integrated prism-based precompensation module for those customers who wish to mitigate the pulse-broadening effects of their microscope optics.
The standard features of the Chameleon remain the same – this overview describes the additional user options provided to access the precomp functions.

## Slide 4: Dispersion: Physical Process
Compensation of chromatic dispersion
  Pulse Length Optimization
  Group Dispersion Delay (GDD)
  Chameleon Vision

## Slide 5: Pulse Length Extention
fs Pulse gets longer by travelling through glass
Pre-compensation – negative Dispersion
Measurement with Autocorrelator CARPE
Target: Creation of a GDD-Curve for individual Setup
Measurement on-site possible

## Slide 6: CARPE - Measurement

## Slide 7: CARPE - Measurement

## Slide 8: GDD - Curves

## Slide 9: Dispersion Settings Menu

## Slide 10: Dispersion Settings: Main Menu
In addition to the standard information, the Chameleon Vision main screen also displays
•The currently-selected precomp ‘curve’
•The current GDD value.
Note that an asterisk (*) beside the GDD value indicates that this value is extrapolated not interpolated.

## Slide 11: Dispersion Settings Menu
New Menu Point: Dispersion Settings

## Slide 12: Dispersion Settings Menu: Select Curve
> Menu Select

## Slide 13: Dispersion Settings Menu: Select Curve
> Menu Select
Displays Current Selected Curve at top of screen.
Nine Available Curves –Menu Up/Down and Select
If MANUAL curve is selected, then the user is prompted to enter a new GDD value, using the rotary knob.
If ZERO curve is selected, then it is the same as Ultra

## Slide 14: Dispersion Settings Menu: New Curve
Use rotary knob to select characters
press knob to accept
Menu Up/Down to go back.
Names can contain characters A –Z ; 1 –9; underscore ( _ ) and full stop (.)
Names can be a maximum 6 characters long.

## Slide 15: Dispersion Settings Menu: Delete Curve
Select curve to be deleted using the rotary knob.
Additional prompts prevent unintentional deletion.

## Slide 16: Dispersion Settings Menu: Calibration
A curve can have a minimum of 3 points, and a maximum of 10.
Inbetween, the GDD value will interpolated
Outside the calibration points, extrapolation gets applied
An asterix (*) shows up:

## Slide 17: Dispersion Settings Menu: Calibration

## Slide 18: Dispersion Settings Menu
A ‘curve’ can be defined with a minimum of 3 points and a maximum of 10.
It is not necessary to cover the entire tuning range 680-1080nm. In this case, if a wavelength is selected which is beyond the currently-selected ‘curve’, a GDD value is extrapolated from the last two known points (up to the maximum available at that wavelength).
It is also possible to select a MANUAL curve and the desired GDD value can be dialled in realtime.

## Slide 19: Dispersion Compensation Module
Usage of „chirpted“ mirrors
  Different wavelengths penetrate with different depths into the dielectric mirror.
  This generated a „Group Delay“
  Several reflections increase / decrease chirp
  The real beam path and setup is not known to Coherent Field Service